June 13, 2025

SPdeS comments Aug 18, 2025

Submersible humidity sensor testing protocol

Major issues from previous tests:

- Probes not thermally equilibrated to water. They are not always submerged long enough to equilibrate.

- Not clear what true temperature and humidity actually is. No reference instrument was used to measure them.

- Probes did not reach same temperature during submerged time, suggesting (a) lack of equilibration, but we’re not sure because (b) there was no reference measurement.

- Adjustment after exposure to air is confusing. This is exacerbated by (a); and (b) would provide clarity about the adjustment.

Link to spreadsheet of previous tests

Dedicate a calibrated stand-alone ASIMET WXT to “live” with the SHS probe and be used to provide references for SHS tests.

Revised protocol (06/13/2025):

- Use an ASIMET WXT as a reference standard for humidity and temperature (and, if outside, at least note wind, sunlight, weather).

- Sync WXT module clock and SHS clock to common standard

- Note time, turn on WXT

- Note time, turn on SHS logger, leave in air for >5 minutes

- Note time, submerge for at least 2 minutes

- Note time, remove from water for at least 30 minutes (at least for next lab test)

  - SPdeS comment: 30 minutes should usually be long enough to equilibrate to wet bulb temperature, but is probably not long enough for the wet sensor to dry. We also need to characterize whether the wet-sock sensor dries off significantly, as this will affect interpretation of the two temperatures. We will be able to do this quicker and with more certainty with the reference measurements from the WXT.

- Repeat (5) and (6) two more times

Data processing

- Produce a netcdf file from the SHS with distinguishable variable names and unique name for the test (e.g. the date).

- Include WXT data in the SHS netcdf file

  - NetCDF files can’t have 2 record dimensions, so 2 files might be better able to accommodate the two separate time dimensions for the SHS and the WXT.

20251208_SHS-test_EGH

Environmental conditions: in ballast tank lab (176 Clark South); fan blowing on WXT and SHS (see picture)

14:43z – WXT logging (ASIWXT203.dat), test tank GPCTD logging (20251208_SHS-test_EGH.txt)

15:00z – start RBR in air

15:09z – submerge probes into tank

15:18z – remove from tank – strap to table (see picture)

16:38z – submerge in tank

16:52z – removed from tank

18:18z – submerge in tank

19:01z – removed from tank – fan turned off

20:35z – submerge in tank

20:56z – remove from tank and let dry overnight